## Datasheet
CALIBRATION DATA JPL STANDARDS LABORATORY
DATE
DESCRIPTION | Plug-in Module | MANUFACTURER |  | Keysight |  |  | BY
MODEL NO. | Select Model | IDENTIFICATION NO. |  |  |  |  | TEMP (°C)
SERIAL NO. |  | INSTRUCTION NO |  | SL0881 |  |  | HUMIDITY (%RH)
Working Standard Information
IDENTIFICATION NO. | INSTRUMENT | MANUFACTURER |  | MODEL NO. | SERIAL NO. | UNCERTAINTY | DUE DATE | NOTES
| Calibrator | Ectron |  | 1140A |  | TUR >= 4.6:1
FALSE | Data Acq Unit | Keysight |  | FALSE | FALSE |  | FALSE
IS033375 | Temp/Humdity | Fluke |  | 2626S | B28385 |  | 2020-06-14 |  | How to Use this Datasheet
| Comments: | [7 empty] | 1. Select UUT Model Number from the drop down list in B6
Test Performed | [8 empty] | 2. Fill in relevant information for UUT and Standards Used
Input Function Tested | UUT Temp Range ( °C ) | Applied Temperature ( °C ) | UUT Measured Value ( °C ) | UUT Error ( °C ) | OEM Tolerance ( +/- °C ) | OEM Minimum Value | OEM Maximum Value | Acceptance Criteria Pass / Fail | 3. For each test point, go to their respective sheet and enter 5 measurements on the far right of the sheet under UUT Reading, next to Test Run. Do Not enter directly into Measured Value on this sheet
Type E (Channel #1) | FALSE | FALSE |  |  | 0 | FALSE | FALSE
|  | -80 |  |  | 1 | -81 | -79 |  | Updating Datasheet
|  | 0 |  |  | 1 | -1 | 1 |  | 1. For DAQ Mainframe: Do Not enter Due Date directly into the cell. Instead, click on the cell and enter the date within the formula in the correct spot
|  | 200 |  |  | 1 | 199 | 201
|  | FALSE |  |  | 1 | -1 | 1 |  | 2. Update Fluke probe as normal
3. Save a blank sheet with updated information
Type J (Channel #6) | FALSE | FALSE |  |  | 0 | FALSE | FALSE
|  | -80 |  |  | 1 | -81 | -79
|  | 0 |  |  | 1 | -1 | 1
|  | 200 |  |  | 1 | 199 | 201
|  | FALSE |  |  | 1 | -1 | 1
Type K (Channel #11) | FALSE | FALSE |  |  | 0 | FALSE | FALSE
|  | -80 |  |  | 80 | FALSE | FALSE
|  | 0 |  |  | 80 | -80 | 80
|  | 200 |  |  | 80 | 120 | 280
|  | 1200 |  |  | 80 | 1120 | 1280
Type T (Channel #17) | FALSE | FALSE |  |  | 0 | FALSE | FALSE
|  | -80 |  |  | 80 | FALSE | FALSE
|  | 0 |  |  | 80 | -80 | 80
|  | 100 |  |  | 80 | 20 | 180
|  | FALSE |  |  | 80 | -80 | 80
Notes:
1 | The user needs to take into account the calibration tolerance for thermocouple wire used during their test to determine the overall tolerance of the temperature measurement.
2 | Maintaining a TUR >= 4.6:1 guarantees that the provided calibration data will have a false accept risk less than 2%, complying with calibration standards such as ANSI Z540.3, and ISO 17025
3

## Total TUR
Test Point | Measurement Value | UUT Accuracy | Measurement Uncertainty | TUR
Type E (0°C) | #DIV/0! | 0 | #DIV/0! | #DIV/0!
Type E (-80°C) | #DIV/0! | 1 | #DIV/0! | #DIV/0!
Type E (0°C) | #DIV/0! | 1 | #DIV/0! | #DIV/0!
Type E (200°C) | #DIV/0! | 1 | #DIV/0! | #DIV/0!
Type E (0°C) | #DIV/0! | 1 | #DIV/0! | #DIV/0!
Type J (0°C) | #DIV/0! | 0 | #DIV/0! | #DIV/0!
Type J (-80°C) | #DIV/0! | 1 | #DIV/0! | #DIV/0!
Type J (0°C) | #DIV/0! | 1 | #DIV/0! | #DIV/0!
Type J (200°C) | #DIV/0! | 1 | #DIV/0! | #DIV/0!
Type J (0°C) | #DIV/0! | 1 | #DIV/0! | #DIV/0!
Type K (0°C) | #DIV/0! | 0 | #DIV/0! | #DIV/0!
Type K (-80°C) | #DIV/0! | 80 | #DIV/0! | #DIV/0!
Type K (0°C) | #DIV/0! | 80 | #DIV/0! | #DIV/0!
Type K (200°C) | #DIV/0! | 80 | #DIV/0! | #DIV/0!
Type K (1200°C) | #DIV/0! | 80 | #DIV/0! | #DIV/0!
Type T (0°C) | #DIV/0! | 0 | #DIV/0! | #DIV/0!
Type T (-80°C) | #DIV/0! | 80 | #DIV/0! | #DIV/0!
Type T (0°C) | #DIV/0! | 80 | #DIV/0! | #DIV/0!
Type T (100°C) | #DIV/0! | 80 | #DIV/0! | #DIV/0!
Type T (0°C) | #DIV/0! | 80 | #DIV/0! | #DIV/0!

## Type E (-200°C)
#DIV/0!
| Error Source | Error Limit | Containment Probability | Uncertainty Type | Distribution | Divisor | Standard Uncertainty | Ci | Ui | Degrees of Freedom | Test Run (n) | UUT Reading (°C)
U1 | εStd_bias_1140A | 0.18 | 95% | B | Normal | 1.96 | 0.091838 | 1 | 0.091838 | ∞
U2 | εStd_resolution_1140A | 0.005 | 100% | B | Uniform | 1.7321 | 0.0028868 | 1 | 0.0028868 | ∞ | 1
U5 | εUUT_resolution | 0.05 | 100% | B | Uniform | 1.7321 | 0.028868 | 1 | 0.028868 | ∞ | 2
U6 | εUUT_repeatability | #DIV/0! | 68% | A | Normal | n/a | #DIV/0! | 1 | #DIV/0! | 4 | 3
UC | Combined Uncertainty: |  | #DIV/0! |  | Error Model for Analysis εX = εStd_bias + εStd_resolution + εUUT_resolution + εUUT_repeatability NOTE: No correlations exist between error sources | [5 empty] | 4
UP | Confidence limits (95%): |  | #DIV/0! | [7 empty] | 5
Effective degrees of freedom |  |  | #DIV/0! | [7 empty] | Average (mean) | #DIV/0!
Standard Bias: |  |  | Uncertainty due to the bias in the Reference Standard.
Standard Resolution: |  |  | Uncertainty due to the digital resultion of the Reference Standard. | [7 empty] | Standard Deviation | #DIV/0!
Repeatability: |  |  | Uncertainty to random variations during the measurement process. This uncertainty in the mean value was estimated by dividing the standard deviation of five test runs by the square root of 5. The measurements were performed under Repeatability Conditions.
[11 empty] | Standard Deviation of Mean | #DIV/0!
UUT Resolution: |  |  | Uncertainty due to the digital resultion of the UUT.
|  |  | t-statistic: | #DIV/0!
|  |  | confidence limit: | #DIV/0!

## Type E (-80°C)
#DIV/0!
| Error Source | Error Limit | Containment Probability | Uncertainty Type | Distribution | Divisor | Standard Uncertainty | Ci | Ui | Degrees of Freedom | Test Run (n) | UUT Reading (°C)
U1 | εStd_bias_1140A | 0.07 | 95% | B | Normal | 1.96 | 0.035715 | 1 | 0.035715 | ∞
U2 | εStd_resolution_1140A | 0.005 | 100% | B | Uniform | 1.7321 | 0.0028868 | 1 | 0.0028868 | ∞ | 1
U5 | εUUT_resolution | 0.05 | 100% | B | Uniform | 1.7321 | 0.028868 | 1 | 0.028868 | ∞ | 2
U6 | εUUT_repeatability | #DIV/0! | 68% | A | Normal | n/a | #DIV/0! | 1 | #DIV/0! | 4 | 3
UC | Combined Uncertainty: |  | #DIV/0! |  | Error Model for Analysis εX = εStd_bias + εStd_resolution + εUUT_resolution + εUUT_repeatability NOTE: No correlations exist between error sources | [5 empty] | 4
UP | Confidence limits (95%): |  | #DIV/0! | [7 empty] | 5
Effective degrees of freedom |  |  | #DIV/0! | [7 empty] | Average (mean) | #DIV/0!
Standard Bias: |  |  | Uncertainty due to the bias in the Reference Standard.
Standard Resolution: |  |  | Uncertainty due to the digital resultion of the Reference Standard. | [7 empty] | Standard Deviation | #DIV/0!
Repeatability: |  |  | Uncertainty to random variations during the measurement process. This uncertainty in the mean value was estimated by dividing the standard deviation of five test runs by the square root of 5. The measurements were performed under Repeatability Conditions.
[11 empty] | Standard Deviation of Mean | #DIV/0!
UUT Resolution: |  |  | Uncertainty due to the digital resultion of the UUT.
|  |  | t-statistic: | #DIV/0!
|  |  | confidence limit: | #DIV/0!

## Type E (0°C)
#DIV/0!
| Error Source | Error Limit | Containment Probability | Uncertainty Type | Distribution | Divisor | Standard Uncertainty | Ci | Ui | Degrees of Freedom | Test Run (n) | UUT Reading (°C)
U1 | εStd_bias_1140A | 0.07 | 95% | B | Normal | 1.96 | 0.035715 | 1 | 0.035715 | ∞
U2 | εStd_resolution_1140A | 0.005 | 100% | B | Uniform | 1.7321 | 0.0028868 | 1 | 0.0028868 | ∞ | 1
U5 | εUUT_resolution | 0.05 | 100% | B | Uniform | 1.7321 | 0.028868 | 1 | 0.028868 | ∞ | 2
U6 | εUUT_repeatability | #DIV/0! | 68% | A | Normal | n/a | #DIV/0! | 1 | #DIV/0! | 4 | 3
UC | Combined Uncertainty: |  | #DIV/0! |  | Error Model for Analysis εX = εStd_bias + εStd_resolution + εUUT_resolution + εUUT_repeatability NOTE: No correlations exist between error sources | [5 empty] | 4
UP | Confidence limits (95%): |  | #DIV/0! | [7 empty] | 5
Effective degrees of freedom |  |  | #DIV/0! | [7 empty] | Average (mean) | #DIV/0!
Standard Bias: |  |  | Uncertainty due to the bias in the Reference Standard.
Standard Resolution: |  |  | Uncertainty due to the digital resultion of the Reference Standard. | [7 empty] | Standard Deviation | #DIV/0!
Repeatability: |  |  | Uncertainty to random variations during the measurement process. This uncertainty in the mean value was estimated by dividing the standard deviation of five test runs by the square root of 5. The measurements were performed under Repeatability Conditions.
[11 empty] | Standard Deviation of Mean | #DIV/0!
UUT Resolution: |  |  | Uncertainty due to the digital resultion of the UUT.
|  |  | t-statistic: | #DIV/0!
|  |  | confidence limit: | #DIV/0!

## Type E (200°C)
#DIV/0!
| Error Source | Error Limit | Containment Probability | Uncertainty Type | Distribution | Divisor | Standard Uncertainty | Ci | Ui | Degrees of Freedom | Test Run (n) | UUT Reading (°C)
U1 | εStd_bias_1140A | 0.06 | 95% | B | Normal | 1.96 | 0.030613 | 1 | 0.030613 | ∞
U2 | εStd_resolution_1140A | 0.005 | 100% | B | Uniform | 1.7321 | 0.0028868 | 1 | 0.0028868 | ∞ | 1
U5 | εUUT_resolution | 0.05 | 100% | B | Uniform | 1.7321 | 0.028868 | 1 | 0.028868 | ∞ | 2
U6 | εUUT_repeatability | #DIV/0! | 68% | A | Normal | n/a | #DIV/0! | 1 | #DIV/0! | 4 | 3
UC | Combined Uncertainty: |  | #DIV/0! |  | Error Model for Analysis εX = εStd_bias + εStd_resolution + εUUT_resolution + εUUT_repeatability NOTE: No correlations exist between error sources | [5 empty] | 4
UP | Confidence limits (95%): |  | #DIV/0! | [7 empty] | 5
Effective degrees of freedom |  |  | #DIV/0! | [7 empty] | Average (mean) | #DIV/0!
Standard Bias: |  |  | Uncertainty due to the bias in the Reference Standard.
Standard Resolution: |  |  | Uncertainty due to the digital resultion of the Reference Standard. | [7 empty] | Standard Deviation | #DIV/0!
Repeatability: |  |  | Uncertainty to random variations during the measurement process. This uncertainty in the mean value was estimated by dividing the standard deviation of five test runs by the square root of 5. The measurements were performed under Repeatability Conditions.
[11 empty] | Standard Deviation of Mean | #DIV/0!
UUT Resolution: |  |  | Uncertainty due to the digital resultion of the UUT.
|  |  | t-statistic: | #DIV/0!
|  |  | confidence limit: | #DIV/0!

## Type E (1000°C)
#DIV/0!
| Error Source | Error Limit | Containment Probability | Uncertainty Type | Distribution | Divisor | Standard Uncertainty | Ci | Ui | Degrees of Freedom | Test Run (n) | UUT Reading (°C)
U1 | εStd_bias_1140A | 0.07 | 95% | B | Normal | 1.96 | 0.035715 | 1 | 0.035715 | ∞
U2 | εStd_resolution_1140A | 0.005 | 100% | B | Uniform | 1.7321 | 0.0028868 | 1 | 0.0028868 | ∞ | 1
U5 | εUUT_resolution | 0.05 | 100% | B | Uniform | 1.7321 | 0.028868 | 1 | 0.028868 | ∞ | 2
U6 | εUUT_repeatability | #DIV/0! | 68% | A | Normal | n/a | #DIV/0! | 1 | #DIV/0! | 4 | 3
UC | Combined Uncertainty: |  | #DIV/0! |  | Error Model for Analysis εX = εStd_bias + εStd_resolution + εUUT_resolution + εUUT_repeatability NOTE: No correlations exist between error sources | [5 empty] | 4
UP | Confidence limits (95%): |  | #DIV/0! | [7 empty] | 5
Effective degrees of freedom |  |  | #DIV/0! | [7 empty] | Average (mean) | #DIV/0!
Standard Bias: |  |  | Uncertainty due to the bias in the Reference Standard.
Standard Resolution: |  |  | Uncertainty due to the digital resultion of the Reference Standard. | [7 empty] | Standard Deviation | #DIV/0!
Repeatability: |  |  | Uncertainty to random variations during the measurement process. This uncertainty in the mean value was estimated by dividing the standard deviation of five test runs by the square root of 5. The measurements were performed under Repeatability Conditions.
[11 empty] | Standard Deviation of Mean | #DIV/0!
UUT Resolution: |  |  | Uncertainty due to the digital resultion of the UUT.
|  |  | t-statistic: | #DIV/0!
|  |  | confidence limit: | #DIV/0!

## Type J (-210°C)
#DIV/0!
| Error Source | Error Limit | Containment Probability | Uncertainty Type | Distribution | Divisor | Standard Uncertainty | Ci | Ui | Degrees of Freedom | Test Run (n) | UUT Reading (°C)
U1 | εStd_bias_1140A | 0.12 | 95% | B | Normal | 1.96 | 0.061226 | 1 | 0.061226 | ∞
U2 | εStd_resolution_1140A | 0.005 | 100% | B | Uniform | 1.7321 | 0.0028868 | 1 | 0.0028868 | ∞ | 1
U5 | εUUT_resolution | 0.05 | 100% | B | Uniform | 1.7321 | 0.028868 | 1 | 0.028868 | ∞ | 2
U6 | εUUT_repeatability | #DIV/0! | 68% | A | Normal | n/a | #DIV/0! | 1 | #DIV/0! | 4 | 3
UC | Combined Uncertainty: |  | #DIV/0! |  | Error Model for Analysis εX = εStd_bias + εStd_resolution + εUUT_resolution + εUUT_repeatability NOTE: No correlations exist between error sources | [5 empty] | 4
UP | Confidence limits (95%): |  | #DIV/0! | [7 empty] | 5
Effective degrees of freedom |  |  | #DIV/0! | [7 empty] | Average (mean) | #DIV/0!
Standard Bias: |  |  | Uncertainty due to the bias in the Reference Standard.
Standard Resolution: |  |  | Uncertainty due to the digital resultion of the Reference Standard. | [7 empty] | Standard Deviation | #DIV/0!
Repeatability: |  |  | Uncertainty to random variations during the measurement process. This uncertainty in the mean value was estimated by dividing the standard deviation of five test runs by the square root of 5. The measurements were performed under Repeatability Conditions.
[11 empty] | Standard Deviation of Mean | #DIV/0!
UUT Resolution: |  |  | Uncertainty due to the digital resultion of the UUT.
|  |  | t-statistic: | #DIV/0!
|  |  | confidence limit: | #DIV/0!

## Type J (-80°C)
#DIV/0!
| Error Source | Error Limit | Containment Probability | Uncertainty Type | Distribution | Divisor | Standard Uncertainty | Ci | Ui | Degrees of Freedom | Test Run (n) | UUT Reading (°C)
U1 | εStd_bias_1140A | 0.08 | 95% | B | Normal | 1.96 | 0.040817 | 1 | 0.040817 | ∞
U2 | εStd_resolution_1140A | 0.005 | 100% | B | Uniform | 1.7321 | 0.0028868 | 1 | 0.0028868 | ∞ | 1
U5 | εUUT_resolution | 0.05 | 100% | B | Uniform | 1.7321 | 0.028868 | 1 | 0.028868 | ∞ | 2
U6 | εUUT_repeatability | #DIV/0! | 68% | A | Normal | n/a | #DIV/0! | 1 | #DIV/0! | 4 | 3
UC | Combined Uncertainty: |  | #DIV/0! |  | Error Model for Analysis εX = εStd_bias + εStd_resolution + εUUT_resolution + εUUT_repeatability NOTE: No correlations exist between error sources | [5 empty] | 4
UP | Confidence limits (95%): |  | #DIV/0! | [7 empty] | 5
Effective degrees of freedom |  |  | #DIV/0! | [7 empty] | Average (mean) | #DIV/0!
Standard Bias: |  |  | Uncertainty due to the bias in the Reference Standard.
Standard Resolution: |  |  | Uncertainty due to the digital resultion of the Reference Standard. | [7 empty] | Standard Deviation | #DIV/0!
Repeatability: |  |  | Uncertainty to random variations during the measurement process. This uncertainty in the mean value was estimated by dividing the standard deviation of five test runs by the square root of 5. The measurements were performed under Repeatability Conditions.
[11 empty] | Standard Deviation of Mean | #DIV/0!
UUT Resolution: |  |  | Uncertainty due to the digital resultion of the UUT.
|  |  | t-statistic: | #DIV/0!
|  |  | confidence limit: | #DIV/0!

## Type J (0°C)
#DIV/0!
| Error Source | Error Limit | Containment Probability | Uncertainty Type | Distribution | Divisor | Standard Uncertainty | Ci | Ui | Degrees of Freedom | Test Run (n) | UUT Reading (°C)
U1 | εStd_bias_1140A | 0.07 | 95% | B | Normal | 1.96 | 0.035715 | 1 | 0.035715 | ∞
U2 | εStd_resolution_1140A | 0.005 | 100% | B | Uniform | 1.7321 | 0.0028868 | 1 | 0.0028868 | ∞ | 1
U5 | εUUT_resolution | 0.05 | 100% | B | Uniform | 1.7321 | 0.028868 | 1 | 0.028868 | ∞ | 2
U6 | εUUT_repeatability | #DIV/0! | 68% | A | Normal | n/a | #DIV/0! | 1 | #DIV/0! | 4 | 3
UC | Combined Uncertainty: |  | #DIV/0! |  | Error Model for Analysis εX = εStd_bias + εStd_resolution + εUUT_resolution + εUUT_repeatability NOTE: No correlations exist between error sources | [5 empty] | 4
UP | Confidence limits (95%): |  | #DIV/0! | [7 empty] | 5
Effective degrees of freedom |  |  | #DIV/0! | [7 empty] | Average (mean) | #DIV/0!
Standard Bias: |  |  | Uncertainty due to the bias in the Reference Standard.
Standard Resolution: |  |  | Uncertainty due to the digital resultion of the Reference Standard. | [7 empty] | Standard Deviation | #DIV/0!
Repeatability: |  |  | Uncertainty to random variations during the measurement process. This uncertainty in the mean value was estimated by dividing the standard deviation of five test runs by the square root of 5. The measurements were performed under Repeatability Conditions.
[11 empty] | Standard Deviation of Mean | #DIV/0!
UUT Resolution: |  |  | Uncertainty due to the digital resultion of the UUT.
|  |  | t-statistic: | #DIV/0!
|  |  | confidence limit: | #DIV/0!

## Type J (200°C)
#DIV/0!
| Error Source | Error Limit | Containment Probability | Uncertainty Type | Distribution | Divisor | Standard Uncertainty | Ci | Ui | Degrees of Freedom | Test Run (n) | UUT Reading (°C)
U1 | εStd_bias_1140A | 0.07 | 95% | B | Normal | 1.96 | 0.035715 | 1 | 0.035715 | ∞
U2 | εStd_resolution_1140A | 0.005 | 100% | B | Uniform | 1.7321 | 0.0028868 | 1 | 0.0028868 | ∞ | 1
U5 | εUUT_resolution | 0.05 | 100% | B | Uniform | 1.7321 | 0.028868 | 1 | 0.028868 | ∞ | 2
U6 | εUUT_repeatability | #DIV/0! | 68% | A | Normal | n/a | #DIV/0! | 1 | #DIV/0! | 4 | 3
UC | Combined Uncertainty: |  | #DIV/0! |  | Error Model for Analysis εX = εStd_bias + εStd_resolution + εUUT_resolution + εUUT_repeatability NOTE: No correlations exist between error sources | [5 empty] | 4
UP | Confidence limits (95%): |  | #DIV/0! | [7 empty] | 5
Effective degrees of freedom |  |  | #DIV/0! | [7 empty] | Average (mean) | #DIV/0!
Standard Bias: |  |  | Uncertainty due to the bias in the Reference Standard.
Standard Resolution: |  |  | Uncertainty due to the digital resultion of the Reference Standard. | [7 empty] | Standard Deviation | #DIV/0!
Repeatability: |  |  | Uncertainty to random variations during the measurement process. This uncertainty in the mean value was estimated by dividing the standard deviation of five test runs by the square root of 5. The measurements were performed under Repeatability Conditions.
[11 empty] | Standard Deviation of Mean | #DIV/0!
UUT Resolution: |  |  | Uncertainty due to the digital resultion of the UUT.
|  |  | t-statistic: | #DIV/0!
|  |  | confidence limit: | #DIV/0!

## Type J (1200°C)
#DIV/0!
| Error Source | Error Limit | Containment Probability | Uncertainty Type | Distribution | Divisor | Standard Uncertainty | Ci | Ui | Degrees of Freedom | Test Run (n) | UUT Reading (°C)
U1 | εStd_bias_1140A | 0.07 | 95% | B | Normal | 1.96 | 0.035715 | 1 | 0.035715 | ∞
U2 | εStd_resolution_1140A | 0.005 | 100% | B | Uniform | 1.7321 | 0.0028868 | 1 | 0.0028868 | ∞ | 1
U5 | εUUT_resolution | 0.05 | 100% | B | Uniform | 1.7321 | 0.028868 | 1 | 0.028868 | ∞ | 2
U6 | εUUT_repeatability | #DIV/0! | 68% | A | Normal | n/a | #DIV/0! | 1 | #DIV/0! | 4 | 3
UC | Combined Uncertainty: |  | #DIV/0! |  | Error Model for Analysis εX = εStd_bias + εStd_resolution + εUUT_resolution + εUUT_repeatability NOTE: No correlations exist between error sources | [5 empty] | 4
UP | Confidence limits (95%): |  | #DIV/0! | [7 empty] | 5
Effective degrees of freedom |  |  | #DIV/0! | [7 empty] | Average (mean) | #DIV/0!
Standard Bias: |  |  | Uncertainty due to the bias in the Reference Standard.
Standard Resolution: |  |  | Uncertainty due to the digital resultion of the Reference Standard. | [7 empty] | Standard Deviation | #DIV/0!
Repeatability: |  |  | Uncertainty to random variations during the measurement process. This uncertainty in the mean value was estimated by dividing the standard deviation of five test runs by the square root of 5. The measurements were performed under Repeatability Conditions.
[11 empty] | Standard Deviation of Mean | #DIV/0!
UUT Resolution: |  |  | Uncertainty due to the digital resultion of the UUT.
|  |  | t-statistic: | #DIV/0!
|  |  | confidence limit: | #DIV/0!

## Type K (-195°C)
#DIV/0!
| Error Source | Error Limit | Containment Probability | Uncertainty Type | Distribution | Divisor | Standard Uncertainty | Ci | Ui | Degrees of Freedom | Test Run (n) | UUT Reading (°C)
U1 | εStd_bias_1140A | 0.12 | 95% | B | Normal | 1.96 | 0.061226 | 1 | 0.061226 | ∞
U2 | εStd_resolution_1140A | 0.005 | 100% | B | Uniform | 1.7321 | 0.0028868 | 1 | 0.0028868 | ∞ | 1
U5 | εUUT_resolution | 0.05 | 100% | B | Uniform | 1.7321 | 0.028868 | 1 | 0.028868 | ∞ | 2
U6 | εUUT_repeatability | #DIV/0! | 68% | A | Normal | n/a | #DIV/0! | 1 | #DIV/0! | 4 | 3
UC | Combined Uncertainty: |  | #DIV/0! |  | Error Model for Analysis εX = εStd_bias + εStd_resolution + εUUT_resolution + εUUT_repeatability NOTE: No correlations exist between error sources | [5 empty] | 4
UP | Confidence limits (95%): |  | #DIV/0! | [7 empty] | 5
Effective degrees of freedom |  |  | #DIV/0! | [7 empty] | Average (mean) | #DIV/0!
Standard Bias: |  |  | Uncertainty due to the bias in the Reference Standard.
Standard Resolution: |  |  | Uncertainty due to the digital resultion of the Reference Standard. | [7 empty] | Standard Deviation | #DIV/0!
Repeatability: |  |  | Uncertainty to random variations during the measurement process. This uncertainty in the mean value was estimated by dividing the standard deviation of five test runs by the square root of 5. The measurements were performed under Repeatability Conditions.
[11 empty] | Standard Deviation of Mean | #DIV/0!
UUT Resolution: |  |  | Uncertainty due to the digital resultion of the UUT.
|  |  | t-statistic: | #DIV/0!
|  |  | confidence limit: | #DIV/0!

## Type K (-80°C)
#DIV/0!
| Error Source | Error Limit | Containment Probability | Uncertainty Type | Distribution | Divisor | Standard Uncertainty | Ci | Ui | Degrees of Freedom | Test Run (n) | UUT Reading (°C)
U1 | εStd_bias_1140A | 0.09 | 95% | B | Normal | 1.96 | 0.045919 | 1 | 0.045919 | ∞
U2 | εStd_resolution_1140A | 0.005 | 100% | B | Uniform | 1.7321 | 0.0028868 | 1 | 0.0028868 | ∞ | 1
U5 | εUUT_resolution | 0.05 | 100% | B | Uniform | 1.7321 | 0.028868 | 1 | 0.028868 | ∞ | 2
U6 | εUUT_repeatability | #DIV/0! | 68% | A | Normal | n/a | #DIV/0! | 1 | #DIV/0! | 4 | 3
UC | Combined Uncertainty: |  | #DIV/0! |  | Error Model for Analysis εX = εStd_bias + εStd_resolution + εUUT_resolution + εUUT_repeatability NOTE: No correlations exist between error sources | [5 empty] | 4
UP | Confidence limits (95%): |  | #DIV/0! | [7 empty] | 5
Effective degrees of freedom |  |  | #DIV/0! | [7 empty] | Average (mean) | #DIV/0!
Standard Bias: |  |  | Uncertainty due to the bias in the Reference Standard.
Standard Resolution: |  |  | Uncertainty due to the digital resultion of the Reference Standard. | [7 empty] | Standard Deviation | #DIV/0!
Repeatability: |  |  | Uncertainty to random variations during the measurement process. This uncertainty in the mean value was estimated by dividing the standard deviation of five test runs by the square root of 5. The measurements were performed under Repeatability Conditions.
[11 empty] | Standard Deviation of Mean | #DIV/0!
UUT Resolution: |  |  | Uncertainty due to the digital resultion of the UUT.
|  |  | t-statistic: | #DIV/0!
|  |  | confidence limit: | #DIV/0!

## Type K (0°C)
#DIV/0!
| Error Source | Error Limit | Containment Probability | Uncertainty Type | Distribution | Divisor | Standard Uncertainty | Ci | Ui | Degrees of Freedom | Test Run (n) | UUT Reading (°C)
U1 | εStd_bias_1140A | 0.07 | 95% | B | Normal | 1.96 | 0.035715 | 1 | 0.035715 | ∞
U2 | εStd_resolution_1140A | 0.005 | 100% | B | Uniform | 1.7321 | 0.0028868 | 1 | 0.0028868 | ∞ | 1
U5 | εUUT_resolution | 0.05 | 100% | B | Uniform | 1.7321 | 0.028868 | 1 | 0.028868 | ∞ | 2
U6 | εUUT_repeatability | #DIV/0! | 68% | A | Normal | n/a | #DIV/0! | 1 | #DIV/0! | 4 | 3
UC | Combined Uncertainty: |  | #DIV/0! |  | Error Model for Analysis εX = εStd_bias + εStd_resolution + εUUT_resolution + εUUT_repeatability NOTE: No correlations exist between error sources | [5 empty] | 4
UP | Confidence limits (95%): |  | #DIV/0! | [7 empty] | 5
Effective degrees of freedom |  |  | #DIV/0! | [7 empty] | Average (mean) | #DIV/0!
Standard Bias: |  |  | Uncertainty due to the bias in the Reference Standard.
Standard Resolution: |  |  | Uncertainty due to the digital resultion of the Reference Standard. | [7 empty] | Standard Deviation | #DIV/0!
Repeatability: |  |  | Uncertainty to random variations during the measurement process. This uncertainty in the mean value was estimated by dividing the standard deviation of five test runs by the square root of 5. The measurements were performed under Repeatability Conditions.
[11 empty] | Standard Deviation of Mean | #DIV/0!
UUT Resolution: |  |  | Uncertainty due to the digital resultion of the UUT.
|  |  | t-statistic: | #DIV/0!
|  |  | confidence limit: | #DIV/0!

## Type K (200°C)
#DIV/0!
| Error Source | Error Limit | Containment Probability | Uncertainty Type | Distribution | Divisor | Standard Uncertainty | Ci | Ui | Degrees of Freedom | Test Run (n) | UUT Reading (°C)
U1 | εStd_bias_1140A | 0.07 | 95% | B | Normal | 1.96 | 0.035715 | 1 | 0.035715 | ∞
U2 | εStd_resolution_1140A | 0.005 | 100% | B | Uniform | 1.7321 | 0.0028868 | 1 | 0.0028868 | ∞ | 1
U5 | εUUT_resolution | 0.05 | 100% | B | Uniform | 1.7321 | 0.028868 | 1 | 0.028868 | ∞ | 2
U6 | εUUT_repeatability | #DIV/0! | 68% | A | Normal | n/a | #DIV/0! | 1 | #DIV/0! | 4 | 3
UC | Combined Uncertainty: |  | #DIV/0! |  | Error Model for Analysis εX = εStd_bias + εStd_resolution + εUUT_resolution + εUUT_repeatability NOTE: No correlations exist between error sources | [5 empty] | 4
UP | Confidence limits (95%): |  | #DIV/0! | [7 empty] | 5
Effective degrees of freedom |  |  | #DIV/0! | [7 empty] | Average (mean) | #DIV/0!
Standard Bias: |  |  | Uncertainty due to the bias in the Reference Standard.
Standard Resolution: |  |  | Uncertainty due to the digital resultion of the Reference Standard. | [7 empty] | Standard Deviation | #DIV/0!
Repeatability: |  |  | Uncertainty to random variations during the measurement process. This uncertainty in the mean value was estimated by dividing the standard deviation of five test runs by the square root of 5. The measurements were performed under Repeatability Conditions.
[11 empty] | Standard Deviation of Mean | #DIV/0!
UUT Resolution: |  |  | Uncertainty due to the digital resultion of the UUT.
|  |  | t-statistic: | #DIV/0!
|  |  | confidence limit: | #DIV/0!

## Type K (1200°C)
#DIV/0!
| Error Source | Error Limit | Containment Probability | Uncertainty Type | Distribution | Divisor | Standard Uncertainty | Ci | Ui | Degrees of Freedom | Test Run (n) | UUT Reading (°C)
U1 | εStd_bias_1140A | 0.08 | 95% | B | Normal | 1.96 | 0.040817 | 1 | 0.040817 | ∞
U2 | εStd_resolution_1140A | 0.005 | 100% | B | Uniform | 1.7321 | 0.0028868 | 1 | 0.0028868 | ∞ | 1
U5 | εUUT_resolution | 0.05 | 100% | B | Uniform | 1.7321 | 0.028868 | 1 | 0.028868 | ∞ | 2
U6 | εUUT_repeatability | #DIV/0! | 68% | A | Normal | n/a | #DIV/0! | 1 | #DIV/0! | 4 | 3
UC | Combined Uncertainty: |  | #DIV/0! |  | Error Model for Analysis εX = εStd_bias + εStd_resolution + εUUT_resolution + εUUT_repeatability NOTE: No correlations exist between error sources | [5 empty] | 4
UP | Confidence limits (95%): |  | #DIV/0! | [7 empty] | 5
Effective degrees of freedom |  |  | #DIV/0! | [7 empty] | Average (mean) | #DIV/0!
Standard Bias: |  |  | Uncertainty due to the bias in the Reference Standard.
Standard Resolution: |  |  | Uncertainty due to the digital resultion of the Reference Standard. | [7 empty] | Standard Deviation | #DIV/0!
Repeatability: |  |  | Uncertainty to random variations during the measurement process. This uncertainty in the mean value was estimated by dividing the standard deviation of five test runs by the square root of 5. The measurements were performed under Repeatability Conditions.
[11 empty] | Standard Deviation of Mean | #DIV/0!
UUT Resolution: |  |  | Uncertainty due to the digital resultion of the UUT.
|  |  | t-statistic: | #DIV/0!
|  |  | confidence limit: | #DIV/0!

## Type T (-200°C)
#DIV/0!
| Error Source | Error Limit | Containment Probability | Uncertainty Type | Distribution | Divisor | Standard Uncertainty | Ci | Ui | Degrees of Freedom | Test Run (n) | UUT Reading (°C)
U1 | εStd_bias_1140A | 0.18 | 95% | B | Normal | 1.96 | 0.091838 | 1 | 0.091838 | ∞
U2 | εStd_resolution_1140A | 0.005 | 100% | B | Uniform | 1.7321 | 0.0028868 | 1 | 0.0028868 | ∞ | 1
U5 | εUUT_resolution | 0.05 | 100% | B | Uniform | 1.7321 | 0.028868 | 1 | 0.028868 | ∞ | 2
U6 | εUUT_repeatability | #DIV/0! | 68% | A | Normal | n/a | #DIV/0! | 1 | #DIV/0! | 4 | 3
UC | Combined Uncertainty: |  | #DIV/0! |  | Error Model for Analysis εX = εStd_bias + εStd_resolution + εUUT_resolution + εUUT_repeatability NOTE: No correlations exist between error sources | [5 empty] | 4
UP | Confidence limits (95%): |  | #DIV/0! | [7 empty] | 5
Effective degrees of freedom |  |  | #DIV/0! | [7 empty] | Average (mean) | #DIV/0!
Standard Bias: |  |  | Uncertainty due to the bias in the Reference Standard.
Standard Resolution: |  |  | Uncertainty due to the digital resultion of the Reference Standard. | [7 empty] | Standard Deviation | #DIV/0!
Repeatability: |  |  | Uncertainty to random variations during the measurement process. This uncertainty in the mean value was estimated by dividing the standard deviation of five test runs by the square root of 5. The measurements were performed under Repeatability Conditions.
[11 empty] | Standard Deviation of Mean | #DIV/0!
UUT Resolution: |  |  | Uncertainty due to the digital resultion of the UUT.
|  |  | t-statistic: | #DIV/0!
|  |  | confidence limit: | #DIV/0!

## Type T (-80°C)
#DIV/0!
| Error Source | Error Limit | Containment Probability | Uncertainty Type | Distribution | Divisor | Standard Uncertainty | Ci | Ui | Degrees of Freedom | Test Run (n) | UUT Reading (°C)
U1 | εStd_bias_1140A | 0.12 | 95% | B | Normal | 1.96 | 0.061226 | 1 | 0.061226 | ∞
U2 | εStd_resolution_1140A | 0.005 | 100% | B | Uniform | 1.7321 | 0.0028868 | 1 | 0.0028868 | ∞ | 1
U5 | εUUT_resolution | 0.05 | 100% | B | Uniform | 1.7321 | 0.028868 | 1 | 0.028868 | ∞ | 2
U6 | εUUT_repeatability | #DIV/0! | 68% | A | Normal | n/a | #DIV/0! | 1 | #DIV/0! | 4 | 3
UC | Combined Uncertainty: |  | #DIV/0! |  | Error Model for Analysis εX = εStd_bias + εStd_resolution + εUUT_resolution + εUUT_repeatability NOTE: No correlations exist between error sources | [5 empty] | 4
UP | Confidence limits (95%): |  | #DIV/0! | [7 empty] | 5
Effective degrees of freedom |  |  | #DIV/0! | [7 empty] | Average (mean) | #DIV/0!
Standard Bias: |  |  | Uncertainty due to the bias in the Reference Standard.
Standard Resolution: |  |  | Uncertainty due to the digital resultion of the Reference Standard. | [7 empty] | Standard Deviation | #DIV/0!
Repeatability: |  |  | Uncertainty to random variations during the measurement process. This uncertainty in the mean value was estimated by dividing the standard deviation of five test runs by the square root of 5. The measurements were performed under Repeatability Conditions.
[11 empty] | Standard Deviation of Mean | #DIV/0!
UUT Resolution: |  |  | Uncertainty due to the digital resultion of the UUT.
|  |  | t-statistic: | #DIV/0!
|  |  | confidence limit: | #DIV/0!

## Type T (0°C)
#DIV/0!
| Error Source | Error Limit | Containment Probability | Uncertainty Type | Distribution | Divisor | Standard Uncertainty | Ci | Ui | Degrees of Freedom | Test Run (n) | UUT Reading (°C)
U1 | εStd_bias_1140A | 0.08 | 95% | B | Normal | 1.96 | 0.040817 | 1 | 0.040817 | ∞
U2 | εStd_resolution_1140A | 0.005 | 100% | B | Uniform | 1.7321 | 0.0028868 | 1 | 0.0028868 | ∞ | 1
U5 | εUUT_resolution | 0.05 | 100% | B | Uniform | 1.7321 | 0.028868 | 1 | 0.028868 | ∞ | 2
U6 | εUUT_repeatability | #DIV/0! | 68% | A | Normal | n/a | #DIV/0! | 1 | #DIV/0! | 4 | 3
UC | Combined Uncertainty: |  | #DIV/0! |  | Error Model for Analysis εX = εStd_bias + εStd_resolution + εUUT_resolution + εUUT_repeatability NOTE: No correlations exist between error sources | [5 empty] | 4
UP | Confidence limits (95%): |  | #DIV/0! | [7 empty] | 5
Effective degrees of freedom |  |  | #DIV/0! | [7 empty] | Average (mean) | #DIV/0!
Standard Bias: |  |  | Uncertainty due to the bias in the Reference Standard.
Standard Resolution: |  |  | Uncertainty due to the digital resultion of the Reference Standard. | [7 empty] | Standard Deviation | #DIV/0!
Repeatability: |  |  | Uncertainty to random variations during the measurement process. This uncertainty in the mean value was estimated by dividing the standard deviation of five test runs by the square root of 5. The measurements were performed under Repeatability Conditions.
[11 empty] | Standard Deviation of Mean | #DIV/0!
UUT Resolution: |  |  | Uncertainty due to the digital resultion of the UUT.
|  |  | t-statistic: | #DIV/0!
|  |  | confidence limit: | #DIV/0!

## Type T (100°C)
#DIV/0!
| Error Source | Error Limit | Containment Probability | Uncertainty Type | Distribution | Divisor | Standard Uncertainty | Ci | Ui | Degrees of Freedom | Test Run (n) | UUT Reading (°C)
U1 | εStd_bias_1140A | 0.07 | 95% | B | Normal | 1.96 | 0.035715 | 1 | 0.035715 | ∞
U2 | εStd_resolution_1140A | 0.005 | 100% | B | Uniform | 1.7321 | 0.0028868 | 1 | 0.0028868 | ∞ | 1
U5 | εUUT_resolution | 0.05 | 100% | B | Uniform | 1.7321 | 0.028868 | 1 | 0.028868 | ∞ | 2
U6 | εUUT_repeatability | #DIV/0! | 68% | A | Normal | n/a | #DIV/0! | 1 | #DIV/0! | 4 | 3
UC | Combined Uncertainty: |  | #DIV/0! |  | Error Model for Analysis εX = εStd_bias + εStd_resolution + εUUT_resolution + εUUT_repeatability NOTE: No correlations exist between error sources | [5 empty] | 4
UP | Confidence limits (95%): |  | #DIV/0! | [7 empty] | 5
Effective degrees of freedom |  |  | #DIV/0! | [7 empty] | Average (mean) | #DIV/0!
Standard Bias: |  |  | Uncertainty due to the bias in the Reference Standard.
Standard Resolution: |  |  | Uncertainty due to the digital resultion of the Reference Standard. | [7 empty] | Standard Deviation | #DIV/0!
Repeatability: |  |  | Uncertainty to random variations during the measurement process. This uncertainty in the mean value was estimated by dividing the standard deviation of five test runs by the square root of 5. The measurements were performed under Repeatability Conditions.
[11 empty] | Standard Deviation of Mean | #DIV/0!
UUT Resolution: |  |  | Uncertainty due to the digital resultion of the UUT.
|  |  | t-statistic: | #DIV/0!
|  |  | confidence limit: | #DIV/0!

## Type T (400°C)
#DIV/0!
| Error Source | Error Limit | Containment Probability | Uncertainty Type | Distribution | Divisor | Standard Uncertainty | Ci | Ui | Degrees of Freedom | Test Run (n) | UUT Reading (°C)
U1 | εStd_bias_1140A | 0.07 | 95% | B | Normal | 1.96 | 0.035715 | 1 | 0.035715 | ∞
U2 | εStd_resolution_1140A | 0.005 | 100% | B | Uniform | 1.7321 | 0.0028868 | 1 | 0.0028868 | ∞ | 1
U5 | εUUT_resolution | 0.05 | 100% | B | Uniform | 1.7321 | 0.028868 | 1 | 0.028868 | ∞ | 2
U6 | εUUT_repeatability | #DIV/0! | 68% | A | Normal | n/a | #DIV/0! | 1 | #DIV/0! | 4 | 3
UC | Combined Uncertainty: |  | #DIV/0! |  | Error Model for Analysis εX = εStd_bias + εStd_resolution + εUUT_resolution + εUUT_repeatability NOTE: No correlations exist between error sources | [5 empty] | 4
UP | Confidence limits (95%): |  | #DIV/0! | [7 empty] | 5
Effective degrees of freedom |  |  | #DIV/0! | [7 empty] | Average (mean) | #DIV/0!
Standard Bias: |  |  | Uncertainty due to the bias in the Reference Standard.
Standard Resolution: |  |  | Uncertainty due to the digital resultion of the Reference Standard. | [7 empty] | Standard Deviation | #DIV/0!
Repeatability: |  |  | Uncertainty to random variations during the measurement process. This uncertainty in the mean value was estimated by dividing the standard deviation of five test runs by the square root of 5. The measurements were performed under Repeatability Conditions.
[11 empty] | Standard Deviation of Mean | #DIV/0!
UUT Resolution: |  |  | Uncertainty due to the digital resultion of the UUT.
|  |  | t-statistic: | #DIV/0!
|  |  | confidence limit: | #DIV/0!

## STD Info
*Unluess otherwise noted, these specifications apply at 23°C ± 3°C, after a 30 minute warm-up period
for one year without calibration.
STD Information |  | Temperature mode:
Manufacturer: |  | Ectron
Model: |  | 1140A
Range: |  | E, J, K, and T (See below)
Resolution: |  | 0.01
Accuracy: |  | See tables below
Type E (°C) |  |  |  |  | Type J (°C)
Temperature Range |  | Error After: |  |  | Temperature Range |  | Error After:
|  | 30 Days | 6 Months | 1 Year |  |  | 30 Days | 6 Months | 1 Year
-270 | <-245 | 0.95 | 1.1 | 1.2 | -210 | <-180 | 0.07 | 0.1 | 0.12
-245 | <-195 | 0.13 | 0.16 | 0.18 | -180 | <-120 | 0.06 | 0.09 | 0.1
-195 | <-155 | 0.07 | 0.09 | 0.1 | -120 | <-50 | 0.06 | 0.07 | 0.08
-155 | <-90 | 0.06 | 0.07 | 0.08 | -50 | <990 | 0.05 | 0.06 | 0.07
-90 | <15 | 0.05 | 0.06 | 0.07 | 990 | 1200 | 0.05 | 0.07 | 0.07
15 | <890 | 0.05 | 0.06 | 0.06
890 | 1000 | 0.05 | 0.06 | 0.07
Type K (°C) |  |  |  |  | Type T (°C)
Temperature Range |  | Error After: |  |  | Temperature Range |  | Error After:
|  | 30 Days | 6 Months | 1 Year |  |  | 30 Days | 6 Months | 1 Year
-270 | <-255 | 1.5 | 1.9 | 2.2 | -270 | <-255 | 1.4 | 1.6 | 1.8
-255 | <-195 | 0.3 | 0.4 | 0.7 | -255 | <-240 | 0.27 | 0.35 | 0.49
-195 | <-115 | 0.1 | 0.11 | 0.12 | -240 | <-210 | 0.17 | 0.24 | 0.3
-115 | <-55 | 0.07 | 0.08 | 0.09 | -210 | <-150 | 0.11 | 0.15 | 0.18
-55 | <1000 | 0.06 | 0.07 | 0.07 | -150 | <-40 | 0.08 | 0.1 | 0.12
1000 | 1372 | 0.06 | 0.08 | 0.08 | -40 | <100 | 0.06 | 0.07 | 0.08
[5 empty] | 100 | 400 | 0.05 | 0.06 | 0.07

## UUT Info
UUT Information | Temperature mode: |  |  | UUT Information | Temperature mode:
Manufacturer: | Keysight |  |  | Manufacturer: | Keysight
Model: | DAQM901A |  |  | Model: | 34901A
Range: | E, J, K, and T |  |  | Range: | E, J, K, and T
Resolution: | 0.01 |  |  | Resolution: | 0.1
Accuracy: | See tables below |  |  | Accuracy: | See tables below
Type E (°C) |  | Type J (°C) |  | Type E (°C) |  | Type J (°C)
Temperature Range | Accuracy | Temperature Range | Accuracy | Temperature Range | Accuracy | Temperature Range | Accuracy
-200°C to -150°C | 1.4 | -210°C to -150°C | 1.6 | -200°C to -150°C | 1.5 | -210°C to -150°C | 1.2
-150°C to 1000°C | 1 | -150°C to 1200°C | 1 | -150°C to 1000°C | 1 | -150°C to 1200°C | 1
Type K (°C) |  | Type T (°C) |  | Type K (°C) |  | Type T (°C)
Temperature Range | Accuracy | Temperature Range | Accuracy | Temperature Range | Accuracy | Temperature Range | Accuracy
-200°C to -100°C | 1.7 | -200°C to -100°C | 1.7 | -200°C to -100°C | 1.5 | -200°C to -100°C | 1.5
-100°C to 1200°C | 0.9 | -100°C to 400°C | 0.9 | -100°C to 1200°C | 1 | -100°C to 400°C | 1